## Sheet1
Directions on Using This Spreadsheet Template:
For Part A: You will only need to enter your actual masses obtained from the laboratory experiment (boxes highlighted in beige). Excel will calculate everything else.
For Part B: You will only need to enter your concentration and ABS values from the UV-Vis instrument (boxes highlighted in beige). Excel will generate a calibration curve.
C-TEACH Academy: Micropipettes |  |  |  | NAME: |  |  |  | % ERROR: If box is "red", you are not pipetting the correct amount based on the stamped volume. Need "BLUE"
%RSD: Repeat deliveries should be close! If not, you are not providing "precise" measurements. Need "GREEN".
Laboratory Check Sheet
Tolerance: +/- the expected volume as stamped on the glassware. |  |  |  |  | 0.997 | g/mL |  | <= Density of Water (be sure to use the correct value, based on the temperature when measurements were performed)
Pipetting 1000 uL | [8 empty] | Making Solutions Via Micropipette
Weight (g) | Theory (g) | Error (g) | Error (mL) |  | % error |  | 100-1000 uL Micropipettes should have an percent error of +/- 2% and precision should be +/- 0.5%
0.995 | 0.997 | 0.002 | 0.002006 |  | 0.38114 |  |  |  |  | [x] | ABS
0.991 | 0.997 | 0.006 | 0.0060181 | [6 empty] | 0 | 0 |  | Note: Regression coefficient (R2) should be a 0.99 or greater. If it is not, one or more of your standards are prepared incorrectly. Remake and try again!
0.993 | 0.997 | 0.004 | 0.004012 | [6 empty] | 0.015 | 0.032
0.995 | 0.997 | 0.002 | 0.002006 | [6 empty] | 0.036 | 0.082
0.992 | 0.997 | 0.005 | 0.005015 | [6 empty] | 0.06 | 0.184
0.9932 | Average | [8 empty] | 0.1125 | 0.258
0.0017889 | STDEV | [8 empty] | 0.15 | 0.312
0.18011 | %RSD
Pipetting 725 uL |  |  |  |  | % error
0.722 | 0.72282 | 0.000825 | 0.00082748 |  | 0.39083 |  | 100-1000 uL Micropipettes should have an percent error of +/- 2% and precision should be +/- 0.5%
0.718 | 0.72282 | 0.004825 | 0.0048395
0.719 | 0.72282 | 0.003825 | 0.0038365
0.721 | 0.72282 | 0.001825 | 0.0018305
0.72 | 0.72282 | 0.002825 | 0.0028335
0.72 | Average
0.0015811 | STDEV
0.2196 | %RSD
Pipetting 515 uL |  |  |  |  | % error
0.514 | 0.51345 | 0.000545 | 0.00054664 |  | 0.40374 |  | 100-1000 uL Micropipettes should have an percent error of +/- 2% and precision should be +/- 0.5%
0.513 | 0.51345 | 0.000455 | 0.00045637
0.511 | 0.51345 | 0.002455 | 0.0024624
0.512 | 0.51345 | 0.001455 | 0.0014594
0.508 | 0.51345 | 0.005455 | 0.0054714
0.5116 | Average
0.0023022 | STDEV
0.44999 | %RSD
Pipetting 208 uL |  |  |  |  | % error
0.205 | 0.20738 | 0.002376 | 0.0023831 |  | 1.4351 |  | 100-1000 uL Micropipettes should have an percent error of +/- 2% and precision should be +/- 0.5%
0.203 | 0.20738 | 0.004376 | 0.0043892
0.205 | 0.20738 | 0.002376 | 0.0023831
0.205 | 0.20738 | 0.002376 | 0.0023831
0.204 | 0.20738 | 0.003376 | 0.0033862
0.2044 | Average
0.00089443 | STDEV
0.43759 | %RSD
Pipetting 100 uL |  |  |  |  | % error |  |  |  | Pipetting 26 uL |  |  |  |  | % error
0.096 | 0.0997 | 0.0037 | 0.0037111 |  | 5.7172 |  | 100-1000 uL Micropipettes should have an percent error of +/- 3% and precision should be +/- 1.0% |  | 0.025 | 0.025922 | 0.000922 | 0.00092477 |  | 5.2203 | 10-100 uL Micropipettes should have an percent error of +/- 3% and precision should be +/- 1.0%
0.089 | 0.0997 | 0.0107 | 0.010732 | [5 empty] | 0.026 | 0.025922 | 7.8e-05 | 7.8235e-05
0.095 | 0.0997 | 0.0047 | 0.0047141 | [5 empty] | 0.025 | 0.025922 | 0.000922 | 0.00092477
0.093 | 0.0997 | 0.0067 | 0.0067202 | [5 empty] | 0.023 | 0.025922 | 0.002922 | 0.0029308
0.097 | 0.0997 | 0.0027 | 0.0027081 | [5 empty] | 0.024 | 0.025922 | 0.001922 | 0.0019278
0.094 | Average | [7 empty] | 0.0246 | Average
0.0031623 | STDEV | [7 empty] | 0.0011402 | STDEV
3.3641 | %RSD | [7 empty] | 4.6349 | %RSD
Pipetting 74 uL |  |  |  |  | % error |  |  |  | Pipetting 10 uL |  |  |  |  | % error
0.073 | 0.073778 | 0.000778 | 0.00078034 |  | 0.90379 |  | 10-100 uL Micropipettes should have an percent error of +/- 3% and precision should be +/- 1.0% |  | 0.01 | 0.00997 | 3e-05 | 3.009e-05 |  | 5.9579 | 10-100 uL Micropipettes should have an percent error of +/- 3% and precision should be +/- 1.0%
0.073 | 0.073778 | 0.000778 | 0.00078034 | [5 empty] | 0.009 | 0.00997 | 0.00097 | 0.00097292
0.073 | 0.073778 | 0.000778 | 0.00078034 | [5 empty] | 0.01 | 0.00997 | 3e-05 | 3.009e-05
0.074 | 0.073778 | 0.000222 | 0.00022267 | [5 empty] | 0.009 | 0.00997 | 0.00097 | 0.00097292
0.073 | 0.073778 | 0.000778 | 0.00078034 | [5 empty] | 0.009 | 0.00997 | 0.00097 | 0.00097292
0.0732 | Average | [7 empty] | 0.0094 | Average
0.00044721 | STDEV | [7 empty] | 0.00054772 | STDEV
0.61095 | %RSD | [7 empty] | 5.8268 | %RSD
This activity has been developed with support from the "Increasing Student Enrollment, Education, and Employment in Chemical Technology" grant (NSF Award 1956274).
C-TEACH Academy: Bridging High School, College, and Industry; Cape Fear Community College - Tracy Holbrook, Program Director of Chemical Technology