## Sheet1
Directions on Using This Spreadsheet Template:
For Part A: You will only need to enter your actual masses obtained from the laboratory experiment (boxes highlighted in beige). Excel will calculate everything else.
For Part B: You will only need to enter your concentration and ABS values from the UV-Vis instrument (boxes highlighted in beige). Excel will generate a calibration curve.
C-TEACH Academy: Volumetric Pipettes |  |  |  | NAME: |  |  |  | Tolerance: If box is "red", you are not pipetting the correct amount based on the stamped volume. Need "BLUE"
STDEV: Repeat deliveries should be close! If not, you are not providing "precise" measurements. Need "GREEN".
Laboratory Check Sheet
Tolerance: +/- the expected volume as stamped on the glassware. |  |  |  |  | 0.99758 | g/mL |  | <= Density of Water (be sure to use the correct value, based on the temperature when measurements were performed)
Pipetting 5.00 mL | [8 empty] | Making Solutions Via Vol. Pipette
Weight (g) | Theory (g) | Error (g) | Error (mL) |  | Tolerance |  | Class A Volumetric Pipettes should have 0.01 tolerance or less
4.98 | 4.9879 | 0.0079 | 0.0079192 |  | 0.0053129 |  |  |  |  | [x] | ABS
4.983 | 4.9879 | 0.0049 | 0.0049119 | [6 empty] | 0 | 0 | Note: Regression coefficient (R2) should be a 0.99 or greater. If it is not, one or more of your standards are prepared incorrectly. Remake and try again!
4.991 | 4.9879 | 0.0031 | 0.0031075 | [6 empty] | 0.015 | 0.032
4.9847 | Average | [8 empty] | 0.036 | 0.082
0.0056862 | STDEV | [8 empty] | 0.06 | 0.184
[10 empty] | 0.1125 | 0.258
[10 empty] | 0.15 | 0.312
Pipetting 10.00 mL |  |  |  |  | Tolerance
9.964 | 9.9758 | 0.0118 | 0.011829 |  | 0.032211 |  | Class A Volumetric Pipettes should have 0.02 tolerance or less
9.944 | 9.9758 | 0.0318 | 0.031877
9.923 | 9.9758 | 0.0528 | 0.052928
9.9437 | Average
0.020502 | STDEV
Pipetting 25.00 mL |  |  |  |  | Tolerance
24.876 | 24.9395 | 0.0635 | 0.063654 |  | 0.10642 |  | Class A Volumetric Pipettes should have 0.03 tolerance or less
24.807 | 24.9395 | 0.1325 | 0.13282
24.817 | 24.9395 | 0.1225 | 0.1228
24.8333 | Average
0.037287 | STDEV
Pipetting 0.700 mL |  |  |  |  | Tolerance
0.604 | 0.69831 | 0.094306 | 0.094535 |  | 0.097876 |  | Class A Volumetric Pipettes should have 0.02 tolerance or less
0.599 | 0.69831 | 0.099306 | 0.099547
0.599 | 0.69831 | 0.099306 | 0.099547
0.60067 | Average
0.0028868 | STDEV
This activity has been developed with support from the "Increasing Student Enrollment, Education, and Employment in Chemical Technology" grant (NSF Award 1956274).
C-TEACH Academy: Bridging High School, College, and Industry; Cape Fear Community College - Tracy Holbrook, Program Director of Chemical Technology